## GENERAL SPEC
SENARAI SPESIFIKASI SEBUT HARGA
SEBUT HARGA MENGGANTIKAN UNIT UTAMA PERALATAN "SIMULTANEOUS THERMAL ANALYSER (TGA/DSC HT)" - SH074/2024/6219800/UPMS
GENERAL SPECIFICATION
I. OBJECTIVES
The specifications provided below are to replace the main unit of the 'Simultaneous Thermal Analyser (TGA/DSC HT)' equipment at the Institute of Nanoscience and Nanotechnology, Universiti Putra Malaysia.
NOTE: 1. Please Tick : [ YES] meets the specifications [ NO ] does not meet the specifications 2. If there are additional accessories please submit a complete description of the specifications separately.
II. | GENERAL SPECIFICATION |  | PLEASE TICK [YES ] or [ NO] (compulsory) | SUPPLIER’S PROPOSAL (please specify)
A. | SUPPLY AND DELIVERY
| a) | This quotation includes supply, delivery, installation, and commissioning on-site at Characterisation Laboratory II, Institute of Nanoscience and Nanotechnology (ION2), Universiti Putra Malaysia, Serdang, Selangor Darul Ehsan.
| b) | The supplier is requested to specify the earliest possible delivery and commissioning date. UPM reserves the right to determine the installation date of the unit.
| c) | UPM reserves the right to extend the period required for delivery and/or completion of the work.
| d) | The supplier will be liable for fulfilling all terms, conditions, and guarantees stipulated in this contract.
| e) | The supplier is requested to demonstrate their competency to successfully execute the job. Past experience in similar projects should be submitted.
| f) | All government taxes, tariffs, regulations, and insurance coverage costs related to the importation, supply, and delivery of the unit shall be the responsibility of the supplier.
| g) | Insurance coverage costs during supply and delivery shall be borne by the supplier.
B. | INSTALLATION AND COMMISSIONING
| a) | The unit supplied shall be delivered on-site by the successful supplier at no extra cost to UPM.
| b) | The successful supplier shall undertake to install and commission the unit supplied on-site by qualified personnel. All demonstrations/testings shall be conducted to assess the performance of the unit supplied, and the data collected shall be analyzed during commissioning. All costs related to this work shall be borne by the supplier.
| c) | If the work carried out in item II.B (b) is not satisfactory, UPM reserves the right not to accept the handover. UPM reserves the right to call upon other professional parties to complete the job satisfactorily, with the cost borne by the supplier.
| d) | All government regulations and procedures must be followed during the installation and commissioning work. Approval from any government authorities, wherever applicable, will be the supplier's responsibility.
C. | MAINTENANCE AND WARRANTY
| a) | The supplier MUST provide at least ONE (1) year of warranty for the supplied unit. The warranty period must begin after complete commissioning.
| b) | The supplier shall outline the details of the maintenance or service plan, which must include performing calibration once a year, for a minimum of TWO (2) years.
| c) | Full maintenance support for the unit supplied shall be available locally. The technical support shall be provided by competent personnel. The supplier shall indicate how and to what extent they will assist in providing the maintenance support.
| d) | The supplier shall be responsible for immediate troubleshooting and/or replacement of any defective or faulty components caused by manufacturing or installation issues. Response time for any breakdown shall be within 48 hours throughout the warranty period.
| e) | Corrective action must be taken within 7 working days. Any delay will require the supplier to extend the warranty period by the number of delayed days, including weekends and public holidays, until the issue is resolved.
D. | OTHER REQUIREMENTS
| a) | The supplier MUST guarantee the availability of good after-sales service for the supplied unit, with a team of service/application support personnel.
| b) | UPM reserves the right not to consider and automatically reject the supplier(s) offer which fails to show any of the above-mentioned items or provides only partial solutions.
| c) | The unit supplied MUST be new. Shipping documents and other relevant documents shall accompany the unit upon delivery. The unit that has been used for demonstration or other purposes will not be accepted.
| d) | A letter and testing report from the manufacturer must be provided.
E. | SITE VISIT
| a) | The supplier is allowed to visit the site for installation by appointment prior to the submission of the Sebutharga document (if necessary).
F. | RESERVED RIGHT
| a) | UPM reserves the right not to consider and automatically reject the supplier(s) offer which fails to comply with any of the above-mentioned items or provides only partial solutions.

## TGA DSC MAIN UNIT
III. SPESIFIKASI TEKNIKAL
1. UNIT UTAMA PERALATAN "SIMULTANEOUS THERMAL ANALYSIS (TGA/DSC)"
SCOPE |  |  | QUANTITY | PLEASE TICK [YES ] or [ NO] (compulsory) | SUPPLIER’S PROPOSAL (please specify)
1 | Technical Requirements
| a. | The specifications provided below are to replace the main unit of the 'Simultaneous Thermal Analyser (TGA/DSC HT)' equipment at the Institute of Nanoscience and Nanotechnology, Universiti Putra Malaysia. | 1
| b. | Supplied Unit to be Replaced Information: Brand: Mettler Toledo Model: TGA/DSC 3+ Furnace: High Temperature Furnace Temperature Range: 22 °C – 1600 °C i. Temperature accuracy: max ±0.5°C or better ii. Temperature precision: max ±0.3°C or better iii. Minimum heating rate: 0.3°C/min or better Sensor: TGA DSC Sensor HT Sensor Types: DTA; DSC; SDTA i. Sensitivity: maximum 0.1 mW or better ii. Temperature resolution: max 0.00003°C or better Balance: XP1 Balance i. Measurement range: ≤1g ii. Resolution: 1.0 µg Data Sampling: 10 data points per second Gas Supply: Gas Controller GC301 Power Option: 240V Operation Manual Language: English Technique: Simultaneous Thermal Analyzer (TGA/DSC); Differential Scanning Calorimetry (DSC); Thermogravimetry (TGA)
| c. | The unit must be compatible with the existing system and software: STARe System TGA/DSC 3+.
| d. | The unit shall be capable of performing applications including, but not limited to, Glass Transition, Heat Capacity (cp), Crystallization, Compositional Analysis, Kinetics, Oxidation Induction Time (OIT), and Temperature Modulated DSC.
2 | Calibration
| a. | The supplier shall perform a full five-point standard calibration during commissioning and provide a calibration certificate.